Profilometer-Bruker Contour GT-I

Basic Operation Instruction

Warning:

- Ensure that there is no object on the scanning stage before setting up the microscope.

- There must be enough space between lens and stage to set up the sample.

- Use the glass plate to prevent dirt on the stage when handling small samples.

- Don’t forget to log off from the FOM after finishing.

Procedure

- Starting

- Measurement Set up

VSI/VXI measurement (Rough sample- 0.3-5nm)

PSI measurement (Smooth sample- step height <0.3nm)

- Looking for fringes and optimize fringes (Focus the sample)

- Measurement and data analysis

- Saving the Data

- Changing the sample

- Finishing

I. Starting

- Open the software .

- Wait for the initialization of the stage and follow the instructions.

Click OK

Click OK

Overview of the screen

- Tool bar- To set up overall system configurations.

- Measurement - To commend the system using measurement parameters.

- Live video- To find the fringes and adjust the fringes.

- Instrument control-To control XY position, Z position, Tip and Tilt, and light intensity.

Note: We can control the XY position, Z position, and Tip and Tilt using control box. We can also use the light control box.

- Turn on the control box. (Don’t turn off the control box after completing the test).

- Set up the sample on the stage.

VSI/VXI measurement (Rough sample- 0.3-5nm)

II. A. Measurement Set up

- Measurement- Choose VSI/VXI

- Start with the objective lens 5x and field of view lens 1.0X.

- Click on Measurement, and choose Measurement Parameters

- Speed-1X

- Back scan and Length – Choose 20um and 50um, respectively. The total movement of the objective lens will be 70um. Depends on sample step height.

- Threshold- Use between 7-15% .

- Illumination- Check Use default (white for VSI measurement)

- Average- Check Average and choose 2 or 3 as you desire.

- Check Enable (End scan 5um after 90%) to save time.

- Processing method- Choose Type VSI

III A. Looking for fringes and optimize fringes.

Note: Fringes are localized around the best focus.

Looking for fringes

- Turn on the auto intensity to see the sample.

- Control the sample position using a control box or mouse on Instrument Control area.

3 options to move the sample using fast, Medium and slow speed.

Z profile is for X direction. Pull up/down the red dot to away/close to the sample.

X profile is for X direction and Y profile is for Y direction. Pull the red dot to X or Y direction.

- Go to edge of the sample to adjust the focus of the image. Can see the fringes when the image is on sharp focus.

- Double click the point that you want to measure. That area will be moved into the center of the crosshair line.

- When you get fringes, change the objective lens and field of view lens to get desired magnification.

Optimize fringes

Note: The white light, Zero order fringes are the highest contracts.

- Use red dot in Tip Tilt to rotate and adjust the number of the fringes.

- Fringes should be perpendicular position. (Spin Red dot in Tip Tilt to get the right position).

- Standard fringes number-3 to 5 Nos.(Pull the red dot up direction to shrink the fringes and down direction to expand the fringes).

- Redo Auto Intensity.

IV. Measurement and data analysis (Term removal, Step height)

- Click on Measurement on Tool bar.

| You can see Surface height, Data and Grayscale Data in Data analyzer (Right side of the screen) - Click on Data. - Choose Term Removal (F-operator) under 3D Filter from Analysis Toolbox (Bottom Right side of the screen) - Choose Step Height under 3D analysis section from Analysis Toolbox |  |
| --- | --- |

- Data Analysis screen will come up under Instrument setting.

Note: Adjust the measurement line using arrow that you want.

1. Edit Mask

- Right click on Term Removal (F-operator).

- Choose Edit settings.

- Term Removal box will come up.

- Choose Tilt only (Plane fit). Click Edit Mask.

- Select the tool shape>select the region> make the region on the image.

Click on Mask, invert Mask then click on OK.

2. Step Height measurement -Click on Step height to get the results.

V. Saving the Data

- Save the file in C:\Program Data\Bruker\Vision\Data\Your Folder

- To save 3D profile file

Click on  button, choose Dataset. Save the data with (*.OPDx) type.

- To save analysis and measurement parameters

Click on Save button, choose Vision Recipe. If the below window will pop up

Click Yes.

- Save in C:\Program Data\Bruker\Vision\Recipes\Vision\Your Folder.

Save the data with (*.visrcp) type.

Recommend: After getting all results, transfer your results to DATA_ Share$ Folder.

Use MMCL server computer in room 532b to get your data by Flash drive. Password (532b).

VI. Changing the sample

- Click on Measurement Setup in the Tool bar.

- Change the Lens 5X.

- Pull up the Red arrow in Z axis using Fast speed and to get enough space between lens and stage to remove the sample and to set up the new sample.

- Change the sample.

- Finishing

- Click on Measurement Setup in the Tool bar.

- Change the Lens 5X.

- Pull up the red arrow in Z axis using Fast speed and to get enough space between lens and stage to remove the sample.

- Make sure save the results before exit the program.

- Exit the program.

- Take the sample and clean the stage.

- Turn off the light box.

- Don’t turn off the control box after completing the test.

- Log out from the FOM.

Log in to MMCL server computer and get your results from Data Share$ Folder.

Analysis Software - Gwydion

PSI measurement (Smooth sample- <0.3nm)

II B. Measurement Set up

- Measurement- Choose PSI

- Start with the objective lens 5x and field of view lens 1.0X.

- Measurement parameters

Threshold- 7%

Illumination- Check Use default (Green for PSI measurement)

Average- Can choose 2 or 3 as you want.

Phase Unwrapping-Choose Standard

III B. Looking for fringes and optimize fringes (Focus the sample)

Looking for fringes

- Turn on the auto intensity to see the sample.

- Control the sample position using a control box or mouse on Instrument Control area.

- Go to edge of the sample to adjust the focus of the image. Can see the fringes when the image is on sharp focus.

- Double click the point that you want to measure. That area will be moved into the center of the crosshair line.

When you get a sharp image, change the objective lens and field of view lens to get desired magnification.

Optimize fringes

Note: Try to get Zero order fringe (White light). In green light, there are many fringes, and all fringes look similar.  It is hard to find Zero order fringe.

- Use red dot in Tip Tilt to rotate and adjust the number of the fringes.

- Use Green light first to find fringes and then change White light to find Zero order fringes.

- Illumination- Remove check Use default and change white, redo auto intensity.

- Then check Use default and choose Green (narrow band), redo auto intensity.

- Fringes should be perpendicular position.  (Spin Red dot in Tip Tilt to get the right position).

- Standard fringes number –1 white (Pull the red dot up direction to shrink the fringes and down direction to expand the fringes).

- Redo Auto Intensity.

Multiple measurement

- Go to Analyze in Menu bar, Click on Setup in Tool bar.

- Database Setup box will come up.

- Click on down arrow in Create and choose Blank.

- Name the file and save with CVS type.

- Save in C:\Program Data\Bruker\Vision\Database\Your Folder. Click on Save.

- Choose Ra, Rq in Basic Stats and Sequence Number and Time stamp in Meta Data.

- Check Overwrite Current Database. Click Save All.

. Click Yes.

- Save in C:\Program Data\Bruker\Vision\Recipes\Vision\Your Folder.

- Save the file with (*.visrcp) type. Click on Save.

- Go to Instrument in Menu bar, choose Measurement Setup in Tool bar.

- Choose Auto loop -3- in Settings under Measurement section.

- Click on Measurement on the Tool bar. Wait for progress.

Click OK. You can see all 3 results.